## Slide 1
6-Axis Force Sensing in Robotic Minimally Invasive Surgery With No Instrument Modification

## Slide 2
Transducer

## Slide 3
Modeling

## Slide 4
Modeling

## Slide 5
Calibration
: The transformation matrices can be lumped into a calibration matrix at point C. Because a calibration matrix can be obtained from every point along the shaft. is fixed for calibration.
Having the normalized signals from the transducer, the calibration matrix can be used to estimate the wrench at point C as:
The modeling equations can be reformatted as below to find the forces and moments at point t.
Additionally:
The parameters for identification are: = 40 parameters

## Slide 6
Calibration
Reference sensor:
ATI Nano43 F/T
Data collection:
Calibration set
Test set
Optimization tool:
MATLAB lsqnonlin
Constrained minimization with the constraints below
  : 0.2 to 10 times the theoretical estimation
  : is a function of the constraints defined on
  : 0.4 to 0.5 m
Multistart for 50 iterations to find the global minimizer

## Slide 7
x
y
z

## Slide 8
Calibration

## Slide 9
Calibration
Friction?

## Slide 10
Calibration

## Slide 11
Limitations
Model Valid
Model Invalid
Model valid
Model Invalid
Model Invalid

## Slide 12
Neural Network Calibration
Reference sensor:
ATI Nano43 F/T
Data collection:
10 cycles during 7 minutes
Optimization tool:
MATLAB fitnet
Training set: first 6 cycles
Validation set: 6-8
Test set: 8-10

## Slide 13
Neural Network Calibration

## Slide 14
Largest error in axial direction
It results in early stop of the training
Neural Network Calibration
Training time < 10 s

## Slide 15
Neural Network Calibration

## Slide 16
Neural Network Calibration (Fz removed)
Training time < 30 s

## Slide 17
Neural Network Calibration (Fz removed)

## Slide 18
Neural Network Calibration (Fz removed)

## Slide 19
Overcoat Test
Wrist Maneuver

## Slide 20
Pinch Test

## Slide 21
Neural Network Calibration (Fz removed, w Gripper Effort)

## Slide 22
Neural Network Calibration (Fz removed, w Gripper Effort)
Training time < 30 s

## Slide 23
Neural Network Calibration (Fz removed, w Gripper Effort)

## Slide 24
Pinch Test

## Slide 25
x
y
z

## Slide 26
 |  | F |  | Train: 1-30 Val:31-40 Test:41-50 |  | S |  | Train: 1-10 Val:11-15 Test:45-50 | 
Layers |  | (10-5)-5 |  |  |  | (10-5)-5 |  |  | 
Model | # Weights | Training Err. | Validation Err. |  | time (s) | Training Err. | Validation Err. |  | time (s)
NDSPE | 475 | 0.0019 | 0.0026 |  | 98it, 3:23 | 0.0017 | 0.0039 |  | 37it, 00:36
NPE | 355 | 0.0019 | 0.0026 |  | 124it, 2:48 | 0.0017 | 0.0037 |  | 63it, 00:40
PE | 235 | 0.0025 | 0.0032 |  | 82it, 1:09 | 0.0022 | 0.0037 |  | 102it, 00:42
Reference sensor:
ATI Nano43 F/T
Data collection:
50 cycles during 70 minutes
Optimization tool:
MATLAB fitnet – Levenberg-Marquardt backpropagation
Networks Trained: categorized based on the inputs to the network
ndspe: normalized-difference-sum-position-effort
npe: normalized-position-effort
pe: position-effort

## Slide 27
NDSPE-F
NPE-F
PE-F
NDSPE-S
NPE-S
PE-S

## Slide 28
NPE-F

## Slide 29
PE-S